DHS-CSAC Jack Rabbit

II FY 2015 JazTM UV-Vis Field Test Report

Jack Rabbit II FY 2015 JazTM UV-Vis Field Test Report

Prepared For:

DHS Chemical Security Analysis Center

Prepared by:

Signature Science, LLC and

Battelle Memorial Institute – Battelle Eastern Science & Technology Center

30 October 2015

DHS-CSAC Jack Rabbit II FY 2015 JazTM UV-Vis Field Test Report

Jack Rabbit II FY 2015 JazTM UV-Vis Field Test Report

Prepared For:

DHS Chemical Security Analysis Center

Prepared by:

Signature Science, LLC and

Battelle Memorial Institute – Battelle Eastern Science & Technology Center

30 October 2015

Table of Contents

1.0	Executive Summary	1

1.1	Background	1

1.2	Results	1

1.3	Recommendations	1

2.0	Background	3

2.1	Trials	3

2.2	JazTM UV-Vis Operation	3

3.0	Results and Conclusions	4

3.1	Trial 1	6

3.2	Trial 2	9

3.3	Trial 3	12

3.4	Trial 4	14

3.5	Trial 5	16

4.0	Quality Assurance and Quality Control	18

4.1	Time Synchronization	18

4.2	Instrument Calibration	18

4.3	Sampling Rate	19

4.4	Environmental Data	19

4.5	Field Test Data	20

4.6	Data Storage and Retrieval	20

5.0	Recommendations	20

5.1	Standard Curve	20

5.2	Time Synchronization	21

Appendix A: Calibration Plots	A-1

Appendix B: Individual Instrument Plots	B-1

Appendix C: Specific Instrument Locations per Trial	C-1

Appendix D: List of Acronyms	D-1

List of Figures

Figure 1. Instrument Layout for Trial 1	7

Figure 2. QC Values for Pre (blue) and Post (red) Release	8

Figure 3. Instrument Layout for Trials 2-5	10

Figure 4. QC Values for Pre (blue) and Post (red) Release	11

Figure 5. QC Values for Pre (blue) and Post (red) Release	13

Figure 6. QC Values for Pre (blue) and Post (red) Release	15

Figure 7. QC Values for Pre (blue) and Post (red) Release	17

List of Tables

Table 1: Overall Test Data	6

Table 2: Data for Trial 1	9

Table 3: Data for Trial 2	12

Table 4: Data for Trial 3	14

Table 5: Data for Trial 4	16

Table 6: Data for Trial 5	18

Table 7: Sampling Rate Data per Instrument for All Trials	19

Table 8: Specific Instrument Locations per Trial	2

1.0	Executive Summary

1.1	Background

Signature Science, LLC (SigSci) is a subcontractor to Battelle Memorial Institute for the Jack Rabbit II program to characterize the behavior of the plume resulting from a large-scale liquid chlorine spill. SigSci also participated in the Jack Rabbit I project which provided data on a large-scale release of chlorine in an outdoor, open environment. For both efforts, SigSci was tasked to operate and maintain custom designed and built vapor sensors based on the Ocean Optic JazTM UV-Vis instrument  to measure chlorine concentration at discrete points throughout the dispersion area.

A series of five releases or trials were carried out at Dugway Proving Grounds (DPG) from August 24-September 3, 2015. These trials ranged in size from approximately five to ten tons per trial. A total of sixteen JazTM UV-Vis instruments were deployed on the release site. Each of the five trials took place in a mock urban layout to simulate the conditions of a trial in a densely populated area. This will aid in the understanding the mechanisms of cloud movement in an urban environment which may have major repercussions on emergency response.

1.2	Results

JazTM UV-Vis instruments successfully recorded data throughout the five chlorine trials with relatively few instrument failures. Recorded chlorine concentrations ranged from below the limit of detection to readings in excess of 100,000 ppm. The accuracy of readings significantly higher than 100,000 cannot be confirmed due to the standards available to establish the standard curve; however, concentrations in excess of 200,000 ppm were measured in the absence of any observable instrument saturation.

Pre- and post-trial calibration checks showed high replicability over the duration of the exercise. This finding highlights that the JazTM UV-Vis instrument is sufficiently precise and rugged to accurately record chlorine gas concentrations after multiple exposures to releases greater than 100,000 ppm.

1.3	Recommendations

The highest standard used to generate the standard curve for each JazTM UV-Vis instrument is 100,000 ppm which precludes confident quantitation at higher concentrations. SigSci recommends that higher concentration standards (>200,000 ppm) are used in the future to establish the upper limit of detection for these instruments. These concentrations will require additional health and safety controls, however, and must be confined to a controlled laboratory setting.

The JazTM UV-Vis instruments experienced variable write-times to the SD memory card which decreased overall temporal resolution. In the future, fitting each instrument with a hard drive or memory device with a much faster write speed will ensure that consistent data is captured across the release. Time synchronization and data retrieval are also complicated by the inability of the JazTM UV-Vis instruments to be wirelessly networked. Future upgrades to allow the instruments to synchronize their internal clocks and transmit data wirelessly will improve data quality and improve ease-of-use for field technicians.

2.0	Background

The Chemical Security Analysis Center (CSAC), part of the Department of Homeland Security (DHS) Science and Technology Directorate, is tasked to study and improve the understanding of rapid large-scale releases of pressurized, liquefied chlorine, a liquefied toxic inhalation hazard (TIH), also known as a toxic industrial chemical (TIC). It is critical to understand the nature of such releases because of the potential for catastrophic accidents, and railcars transporting TIH materials could be attractive targets to terrorists. A containment breach in a railcar transporting these materials, and the subsequent release of toxic gases, has the potential to cause significant numbers of fatalities and injuries.

Signature Science, LLC (SigSci) is a subcontractor to Battelle Memorial Institute for the Jack Rabbit II program to characterize the behavior of the plume resulting from a large-scale liquid chlorine spill. SigSci also participated in the 2010 Jack Rabbit I project which provided data on a large-scale release of chlorine in an outdoor, open environment. For both efforts, SigSci was tasked to operate and maintain Ocean Optic JazTM UV-Vis sensors to measure chlorine concentration at discrete points throughout the dispersion area.

2.1	Trials

A series of five trials were carried out at Dugway Proving Grounds (DPG) from August 24th to September 3rd, 2015. These ranged from approximately five to ten tons per trial. A total of sixteen JazTM UV-Vis instruments were deployed throughout the release site. The JazTM UV-Vis instruments were deployed relatively close to the release site to capitalize on their previously demonstrated ability to measure chlorine concentrations up to 100,000 ppm. All instruments were placed on the ground with their respective air inlets approximately 0.3 meters above ground level.

Each of the five trials took place in a mock urban layout to simulate the conditions of a release in a densely populated area. Very little data is available on the behavior of large scale releases in such confined urban environments. This includes not only the behavior and concentration of gases between buildings but also the intake of gas into the buildings and indoor environments themselves (due to windows, ventilation, etc.). Understanding the mechanisms of cloud movement in an urban environment as well as the amount of gas taken into the structures themselves may have major repercussions on emergency response.

2.2	JazTM UV-Vis Operation

The JazTM UV-Vis instruments were powered and operated in stand-alone mode. Stand-alone mode required daily set up prior to each trial. Each instrument autonomously monitored and stored the measurement data during the trial. Two SigSci personnel were on site during the trials to maintain and calibrate the instruments and to download data.

The JazTM UV-Vis instruments were calibrated with chlorine gas using four standard concentrations using ambient air as the  zero point (0, 100, 1,000, 10,000, and 100,000 ppm) at SigSci in Austin, TX prior to field deployment. This recorded calibration also verified sensor operation and sensitivity (Appendix A). After arriving at DPG, the instruments were challenged with all four standards during the first day of setup. Each day before a test trial, each instrument was challenged with ambient (zero) air followed by the 10,000 ppm standard contained in a Tedlar® bag. A second quality control (QC) check or challenge was performed each day following the trial to verify the calibration remained valid.  The JazTM UV-Vis test plan called for the instruments to be recalibrated if the 10,000 ppm challenge (QC check) was more than 10% off from the calibration curve. Given the 30 minute travel distance from the test site to V Grid (the JazTM UV-Vis prep location)  system recalibration was not practical and was not performed, though pre-trial and post-trial QC checks were recorded.  During the instrument set up prior to each trial, the on-board Jaz data processor (DPU) was cycled through the startup procedure and the clock time checked and set to local time with a clock timer synchronized to the JR 2 command center master clock. The QC checks (challenges) were performed once the instruments were thermally stable, i.e. the heated spectrometer had reached 45oC.

During each test, approximately two-second resolution data was saved onto an internal SD card. The detection limits are a function of data resolution as longer averaging times are required for lower detection limits. The calibration range for each instrument was 0 to 100,000 ppm.  The absorbance values between 0 and 100 ppm for each instrument were essentially the same. The baseline for the instruments varied from 200 to 1000 ppm for instruments 001 and 015.  For these two reasons, we consider the detection limit to be 1000 ppm though lower numbers were recorded, during the trials

Following each trial, SigSci personnel retrieved the data from the JazTM UV-Vis instrument’s SD card and downloaded the data onto a laptop computer. Data retrieved from the JazTM UV-VIS instruments are in the form of comma-delimited text files. These files were imported into Excel spreadsheets and scanned for “hits” of the target compounds by unit and location. Data from units showing detectable concentrations were plotted to show concentration vs. time across the test period (Appendix B).  The text files with the raw data  and instrument location files for each trial were delivered to Mr. Kendal Ferguson at DPG via AMRDEC 24 to 48 hours after each trial.  Within 48 hours of each trial the excel data files and a short preliminary summary were delivered to Dr. Shannon Fox by AMRDEC.

3.0	Results and Conclusions

A total of five trials took place within the mock urban layout. Sixteen JazTM UV-Vis instruments were deployed for each of these trials. A brief overview of the trials is shown in Table 1.  The sixteen instruments were deployed as described in the JR II Mock Urban Plan v 9.2 with two deviations.  The first deviation was no  JazTM UV-Vis instrument was placed on the top of the three high and two wide CONEX stack (CONEX 11.4).  The reason was SigSci personnel needed to access the instruments up to 3 times per trial.  Neither of these two persons had requisite training to travel on the man lift and since the only means to get to the top of the stack was via a man lift, a Jaz instrument was not placed on the top of CONEX 11.4.  For Trial 1 no instrument was placed  directly downwind of CONEX 11.4.   For Trials 2-5 instrument 002 was placed downwind and against the base of the stack where the two bottom CONEXs met.  The second deviation from the Mock Urban Plan was placing an instrument (006) downwind of  CONEX 2.1 for Trial 1. This location was at the suggestion of Dr. Fox and Dr. Tom Spicer.   After Trail 1, at the request of  Dr. Fox,  instrument 006 was relocated  upwind of the CONEX grid at survey point P10 and remained there for Trials 2-5.

For Trial 1 (Figure1) the locations had no survey markings.  Prior to Trial 2, DPG personnel provided survey locations indicated by a J followed by a number that corresponded to the CONEX upwind of the instrument  or in other words the instrument was placed on the downwind side of the indicated CONEX.  Four instruments , 008, 014, 015 and 016, were located on the upwind side of the indicated CONEX. These four instruments were in support of the Lawrence Berkley National Lab’s (LBL) indoor studies and were set up at the inlet to the indicated CONEX.  For Trial 1 these four instruments were at individual CONEX inlets.  For Trials 2-5 the four instruments were repositioned per the request  from Dr. Mike Sohn (LBL) and Dr. Fox.  Instruments 008 and 014 were located on each side of the inlet to CONEX 11.5 and instruments 015 and 016 were on each side of the inlet to CONEX 11.3.   Location details can be found in Appendix C.  Furthermore, Dr. Fox used the results from Trial 1 to relocate other JazTM UV-Vis instruments for Trials 2-5 as shown in Figure 3.

All data was captured as parts per million (ppm) and was not adjusted to compensate for temperature or pressure. We assumed that the linear range for accurate quantitation of chlorine concentration extended 15% above the highest calibration standard used. Therefore, measurements up to 115,000 ppm are considered to be accurate and quantitative. Accuracy above this range cannot be verified; however, measurements were made well above this range prior to any evidence of saturation. Saturation was defined as the inability of an instrument to record an absorbance value indicated by a “nan” in the data spreadsheet’s absorbance column.

We previously reported that several instruments reached saturation as part of the preliminary data delivered to DHS-CSAC immediately following the conclusion of each trial. This was based on measurements over a specific absorbance threshold. Upon more thorough data review, these instruments did not appear to reach true saturation (with a single exception) although these measurements were well beyond the linear range for quantitation as described above.

Baseline or background measurements for each instrument were calculated by averaging the readings taken during the last full minute preceding the trial. Baseline values typically returned to within one standard deviation of this average within 17 minutes after the start of the trial. Plots of the chlorine concentration (ppm) over these 17 minutes for each instrument are shown in Appendix B. For Appendix B, plots labeled with red on the y-axis (concentration) denote a different scale (0 to 250,000 ppm) than those labeled in black (0 to 100,000 ppm). JazTM UV-Vis measurements in excess of 100,000 ppm were relatively infrequent across the five trials.

Table 1: Overall Test Data

| Trial Date | Tonnage/Lbs. | Jaz Instruments | Instruments Functioning | Instruments Disabled | Instruments Saturated | Clock Synchronization |
| --- | --- | --- | --- | --- | --- | --- |
| 8/24/2015 | 5 Ton/ 9940 Lbs | 16 | 13 | 3 | 0 | Within 60s |
| 8/28/2015 | 8 Ton/ 17969 Lbs | 16 | 12 | 4 | 0 | Within 1s |
| 8/29/2015 | 5 Ton/ 9947 Lbs | 16 | 15 | 1 | 0 | Within 1s |
| 9/1/2015 | 7.7 Ton/ 15366 Lbs | 16 | 16 | 0 | 1 | Within 1s |
| 9/3/2015 | 9 Ton/ 18306 Lbs | 16 | 15 | 1 | 0 | Within 1s |

3.1	Trial 1

Trial 1 was carried out on August 24, 2015. Instruments were placed within the mock urban layout as illustrated by Figure 1. More specific instrument locations including proximity to survey points are located in Appendix C.

Figure 1. Instrument Layout for Trial 1

Quality control checks were performed on each instrument before and after the trial using a 10,000 ppm chlorine standard. Figure 2 shows high concordance between the pre-trial (blue columns) and post-trial (red columns) values. The missing post trial (red) column indicates an instrument that was disabled and did not collect data.

For 12 instruments the post-trial QC checks were lower than the pre-trial checks though the differences were minor. Both checks for instrument 011 were approximately 40% less than the standard concentration and less than the other instruments.   The reason for the lower detection was not determined.  In subsequent trials instrument 011 had QC checks at or above 8000 ppm.  For all the Trials, a Tedlar® bag was filled with fresh 10,000 ppm standard prior to the trial set up.  The same bag was refreshed with standard prior to the post trial check and data retrieval.   The QC data suggests that the release event, which in some cases measured significantly higher than 100,000 ppm, did not damage or otherwise impair the function or accuracy of the JazTM UV-Vis instrumentation.

Figure 2. QC Values for Pre (blue) and Post (red) Trial

As part of Trial 1, thirteen instruments successfully recorded the trial. Two of these instruments registered readings significantly above the highest calibration standard (100,000 ppm) and as such these readings cannot be confirmed to be accurate.  Instrument 006 measured 234,402 ppm at an absorbance of 3.75 was initially considered to have saturated. After a review of the data plot and the fact the  instrument did record an absorbance it was determined that saturation may not have occurred.  The remaining instruments registered peak readings which fell along the linear portion of the standard curve and can therefore be considered quantitative. For the first trial, internal instrument clocks were synchronized to within 60 seconds using mobile phone time i.e. not the official DPG JR2 control time keeping system. Instruments were synchronized to official DPG time within +/- 1 second for all remaining trials. Three instruments  (010, 012 and 015) did not collect data.  All three lost power sometime after the pre-trial QC check and before the release. The reason for the lost was later determined to be from loose electrical connections which were subsequently crimped to tighten the connection.

Table 2: Data for Trial 1

| Trial Date | Tonnage/Lbs. |  | Instruments Functioning |  | Instruments Disabled |  | Instruments Saturated | Clock Synchronization |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8/24/2015 | 5 Ton/ 9940 Lbs |  | 13 |  | 3 |  | 0 | Within 60s |  |
| Instrument |  | Baseline (ppm) |  | Highest Measurement (ppm) |  | Evidence of Saturation? |  |  | Quantitative Value? |
| JAZ-001 |  | 924 |  | 101,226 |  | N |  |  | Y |
| JAZ-002 |  | 452 |  | 126,291 |  | N |  |  | N |
| JAZ-003 |  | 874 |  | 38,940 |  | N |  |  | Y |
| JAZ-004 |  | 232 |  | 50,163 |  | N |  |  | Y |
| JAZ-005 |  | 698 |  | 27,154 |  | N |  |  | Y |
| JAZ-006 |  | 444 |  | 234,402 |  | N |  |  | N |
| JAZ-007 |  | 887 |  | 88,556 |  | N |  |  | Y |
| JAZ-008 |  | 261 |  | 5,081 |  | N |  |  | Y |
| JAZ-009 |  | 172 |  | 81,836 |  | N |  |  | Y |
| JAZ-010 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-011 |  | 300 |  | 63,794 |  | N |  |  | Y |
| JAZ-012 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-013 |  | 364 |  | 18,700 |  | N |  |  | Y |
| JAZ-014 |  | 231 |  | 11,190 |  | N |  |  | Y |
| JAZ-015 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-016 |  | 577 |  | 14,086 |  | N |  |  | Y |

3.2	Trial 2

Trial 2 was carried out on August 28, 2015. Instruments were placed within the mock urban layout as illustrated by Figure 3. These instrument locations remained consistent through Trial 5. More specific instrument locations including proximity to survey points are located in Appendix C.

Figure 3. Instrument Layout for Trials 2-5

Quality control checks were performed on each instrument before and after the trial using the 10,000 ppm chlorine standard. Figure 4 shows high concordance between the pre-trial (blue columns) and post-trial (red columns) values. The missing post-trial (red) column indicates that three instruments became disabled after the pre-trial check.  Instrument 010 was disabled before the pre-trial check. Details on the disabled instruments can be found in the next paragraph. All instruments which recorded both the pre- and post-trial QC check registered a minor decrease in the post-trial QC value suggesting that the total chlorine concentration of the standard dropped slightly over the course of the day. For the instruments that collected data, the pre-trial QC checks were within 20% of the 10,000 ppm standard.  The QC data suggests that the instruments continued to perform within expected specifications throughout each progressive release event

Figure 4. QC Values for Pre (blue) and Post (red) Trial

As part of Trial 2, twelve instruments successfully recorded the release. For this as well as subsequent trials, the internal instrument clocks were synchronized to within one second of the official DPG (JR2 control) test clock. Instruments 006 and 012 did not register a significant increase throughout the trial. Pre-trial and post-trial QC checks indicate both instruments were properly collecting data. The highest recorded chlorine concentration for any of the instruments was 95,509 ppm; therefore all measurements for this trial fell within the linear portion of the standard curve and can be considered quantitative.  Four instruments (001, 007, 010 and 015) did not collect.  Instruments 001, 007 and 015 failed due to a power loss after the pre trial QC check and before the release. As for Trial 1, crimping/recrimping the power leads fixed this problem,  Instrument 010 had a light source failure. The light source was replaced and the instrument recalibrated prior to Trial 3.

Table 3: Data for Trial 2

| Trial Date | Tonnage/Lbs. |  | Instruments Functioning |  | Instruments Disabled |  | Instruments Saturated | Clock Synchronization |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8/28/2015 | 8 Ton/ 17969 Lbs |  | 12 |  | 4 |  | 0 | Within 1s |  |
| Instrument |  | Baseline (ppm) |  | Highest Measurement (ppm) |  | Evidence of Saturation? |  |  | Quantitative Value? |
| JAZ-001 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-002 |  | 473 |  | 20,634 |  | N |  |  | Y |
| JAZ-003 |  | 943 |  | 35,224 |  | N |  |  | Y |
| JAZ-004 |  | 271 |  | 34,761 |  | N |  |  | Y |
| JAZ-005 |  | 706 |  | 47,575 |  | N |  |  | Y |
| JAZ-006 |  | 554 |  | 601 |  | N |  |  | Y |
| JAZ-007 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-008 |  | 165 |  | 27,387 |  | N |  |  | Y |
| JAZ-009 |  | 155 |  | 95,509 |  | N |  |  | Y |
| JAZ-010 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-011 |  | 280 |  | 52,742 |  | N |  |  | Y |
| JAZ-012 |  | 347 |  | 379 |  | N |  |  | Y |
| JAZ-013 |  | 360 |  | 34,124 |  | N |  |  | Y |
| JAZ-014 |  | 217 |  | 26,826 |  | N |  |  | Y |
| JAZ-015 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-016 |  | 469 |  | 20,408 |  | N |  |  | Y |

3.3	Trial 3

Trial 3 was carried out on August 29, 2015.  Instruments were placed within the mock urban layout as illustrated by Figure 3. More specific instrument locations including proximity to survey points are located in Appendix C.

As with each trial, quality control checks were performed on each instrument before and after the trial using the 10,000 ppm chlorine standard. Figure 5 continues to show high concordance between the pre-trial (blue columns) and post-trial (red columns) values. The missing post-trial check for instrument 010 indicates it failed to collect data after the pre-trial check.  The QC checks for 13 instruments were within 20% of the standard. Four instruments (001, 002, 006 and 015) had post-trial checks higher than the pre-trial check. The QC data suggests that the instruments continued to perform within expected specifications throughout each progressive release event.

Figure 5. QC Values for Pre (blue) and Post (red) Trial

Fifteen instruments successfully recorded the third trial. The highest recorded chlorine concentration for any of the instruments was 49,754 ppm; therefore all measurements for this release fell within the linear portion of the standard curve and can be considered quantitative.

Instruments  006, 011 and 012  did not register a significant increase above background throughout the trial. Pre-trial and post-trial QC checks indicate the instruments were properly collecting data.  Instrument 010 failed to record data due to a bad SD card. The card was replaced for Trail 4.

Table 4: Data for Trial 3

| Trial Date | Tonnage/Lbs. |  | Instruments Functioning |  | Instruments Disabled |  | Instruments Saturated | Clock Synchronization |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8/29/2015 | 5 Ton/ 9947 Lbs |  | 15 |  | 1 |  | 0 | Within 1s |  |
| Instrument |  | Baseline (ppm) |  | Highest Measurement (ppm) |  | Evidence of Saturation? |  |  | Quantitative Value? |
| JAZ-001 |  | 937 |  | 10,863 |  | N |  |  | Y |
| JAZ-002 |  | 433 |  | 5,652 |  | N |  |  | Y |
| JAZ-003 |  | 908 |  | 22,278 |  | N |  |  | Y |
| JAZ-004 |  | 221 |  | 23,254 |  | N |  |  | Y |
| JAZ-005 |  | 712 |  | 27,308 |  | N |  |  | Y |
| JAZ-006 |  | 423 |  | 462 |  | N |  |  | Y |
| JAZ-007 |  | 880 |  | 20,202 |  | N |  |  | Y |
| JAZ-008 |  | 334 |  | 9,241 |  | N |  |  | Y |
| JAZ-009 |  | 150 |  | 49,754 |  | N |  |  | Y |
| JAZ-010 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-011 |  | 291 |  | 300 |  | N |  |  | Y |
| JAZ-012 |  | 332 |  | 365 |  | N |  |  | Y |
| JAZ-013 |  | 312 |  | 17,610 |  | N |  |  | Y |
| JAZ-014 |  | 256 |  | 8,875 |  | N |  |  | Y |
| JAZ-015 |  | 1016 |  | 12,693 |  | N |  |  | Y |
| JAZ-016 |  | 473 |  | 12,644 |  | N |  |  | Y |

3.4	Trial 4

Trial 4 was carried out on September 1, 2015. Instruments were placed within the mock urban layout as illustrated by Figure 3. More specific instrument locations including proximity to survey points are located in Appendix C.

As with each trial, quality control checks were performed on each instrument before and after the trial using the 10,000 ppm chlorine standard. Figure 6 continues to show high concordance between the pre-trial (blue columns) and post-trial (red columns) values.  As shown in Figure 6 the pre-trial and post-trial QC checks remained consistent.  All sixteen instruments were within 20% of the standard, with fifteen instruments recording minor decreases in concentration from the pre- trial to the post-trial checks. Instrument 010 appeared to have an increase in sensitivity with both QC checks recording values above the standard.  The reason for this increase was not determined. The QC data suggests that the instruments continued to perform within expected specifications throughout each progressive release event.

Figure 6. QC Values for Pre (blue) and Post (red) Trial

All sixteen instruments successfully operated during the fourth trial. The highest recorded chlorine concentrations for instruments 009 and 011 fell above the confirmed linear portion of the standard curve and cannot be confirmed to be quantitative measures.  In addition instrument 009 was not able to record an absorbance value and was considered to have saturated for approximately 90 seconds. The 210,000 ppm in Table 5 should be considered the saturation point though the value is arbitrary. This concentration corresponds to a “nan” absorbance in the data spreadsheet and is slightly above the 209,240 recorded at the highest absorbance of 4.46.  The remainder of instruments fell within the linear range throughout the release.

Instrument 006 again did not register a significant increase above background throughout the trial. The instrument operated correctly as part of the pre-trial and post-trial QC checks which suggests the instrument was shielded by CONEX 1.1.

Table 5: Data for Trial 4

| Trial Date | Tonnage/Lbs. |  | Instruments Functioning |  | Instruments Disabled |  | Instruments Saturated | Clock Synchronization |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 9/1/2015 | 7.7 Ton/ 15366 Lbs |  | 16 |  | 0 |  | 1 | Within 1s |  |
| Instrument |  | Baseline (ppm) |  | Highest Measurement (ppm) |  | Evidence of Saturation? |  |  | Quantitative Value? |
| JAZ-001 |  | 961 |  | 29,072 |  | N |  |  | Y |
| JAZ-002 |  | 454 |  | 16,431 |  | N |  |  | Y |
| JAZ-003 |  | 885 |  | 49,996 |  | N |  |  | Y |
| JAZ-004 |  | 197 |  | 65,534 |  | N |  |  | Y |
| JAZ-005 |  | 730 |  | 52,047 |  | N |  |  | Y |
| JAZ-006 |  | 384 |  | 432 |  | N |  |  | Y |
| JAZ-007 |  | 810 |  | 25,171 |  | N |  |  | Y |
| JAZ-008 |  | 274 |  | 28,758 |  | N |  |  | Y |
| JAZ-009 |  | 195 |  | 210,000 |  | Y |  |  | N |
| JAZ-010 |  | 507 |  | 8,729 |  | N |  |  | Y |
| JAZ-011 |  | 257 |  | 172,000 |  | N* |  |  | N |
| JAZ-012 |  | 384 |  | 52,808 |  | N |  |  | Y |
| JAZ-013 |  | 348 |  | 32,871 |  | N |  |  | Y |
| JAZ-014 |  | 198 |  | 20,410 |  | N |  |  | Y |
| JAZ-015 |  | 1,016 |  | 26,148 |  | N |  |  | Y |
| JAZ-016 |  | 482 |  | 20,501 |  | N |  |  | Y |

* One data point without a measured absorbance

3.5	Trial 5

Trial 5 was carried out on September 3, 2015. Instruments were placed within the mock urban layout as illustrated by Figure 3. More specific instrument locations including proximity to survey points are located in Appendix C.

As with each trial, quality control checks were performed on each instrument before and after the trial using the 10,000 ppm chlorine standard. Figure 7 continues to show high concordance between the pre-trial (blue columns) and post-trial (red columns) values.  The missing columns for instrument 006 indicate it was disabled during this trial. Details on 006 can be found in the next paragraph. As with the other four trials, the pre-trial and post-trial checks were consistent. As in Trial 4 instrument 010 appeared to have an increase in sensitivity with both checks above the standard concentration. The QC data suggests that the instruments continued to perform within expected specifications throughout the release event.

Figure 7. QC Values for Pre (blue) and Post (red) Trial

Fifteen instruments successfully recorded the fifth and final trial. The highest recorded chlorine concentrations for instruments 004, 009 and 011 fell above the confirmed linear portion of the standard curve and cannot be confirmed to be quantitative measures. Instrument 011 had the highest recorded concentration, 239,785 ppm at a measured absorbance of 5.22. This is in contrast to instrument 009 in Trial 4 that was considered saturated at 210,000 ppm when the instrument could not record an absorbance value. These high concentrations underscore the recommendation that more instrument testing needs to be performed to determine the maximum absorbance each instrument can measure accurately and with precision.  Instrument 005 whose highest measurement was 110,090 ppm fell within 15% of the highest standard used to construct the standard curve and is thereby presumed to represent an accurate quantitative measure. The remainder of instruments fell within the linear range throughout the trial.  Instrument 006 failed to initialize when connected to the battery. The problem was traced to the DPU (data processing unit) which would not cycle.  Without sufficient time to enable repairs the instrument was not in service for the trial.

Table 6: Data for Trial 5

| Trial Date | Tonnage/Lbs. |  | Instruments Functioning |  | Instruments Disabled |  | Instruments Saturated | Clock Synchronization |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 9/3/2015 | 9 Ton/ 18306 Lbs |  | 15 |  | 1 |  | 0 | Within 1s |  |
| Instrument |  | Baseline (ppm) |  | Highest Measurement (ppm) |  | Evidence of Saturation? |  |  | Quantitative Value? |
| JAZ-001 |  | 851 |  | 36,050 |  | N |  |  | Y |
| JAZ-002 |  | 415 |  | 21,027 |  | N |  |  | Y |
| JAZ-003 |  | 876 |  | 61,949 |  | N |  |  | Y |
| JAZ-004 |  | 333 |  | 135,061 |  | N |  |  | N |
| JAZ-005 |  | 734 |  | 110,090 |  | N |  |  | Y |
| JAZ-006 |  | Disabled |  |  |  |  |  |  |  |
| JAZ-007 |  | 868 |  | 41,981 |  | N |  |  | Y |
| JAZ-008 |  | 347 |  | 32,570 |  | N |  |  | Y |
| JAZ-009 |  | 107 |  | 127,844 |  | N |  |  | N |
| JAZ-010 |  | 583 |  | 14,762 |  | N |  |  | Y |
| JAZ-011 |  | 249 |  | 239,785 |  | N* |  |  | N |
| JAZ-012 |  | 547 |  | 67,287 |  | N |  |  | Y |
| JAZ-013 |  | 293 |  | 43,057 |  | N |  |  | Y |
| JAZ-014 |  | 272 |  | 27,373 |  | N |  |  | Y |
| JAZ-015 |  | 967 |  | 38,444 |  | N |  |  | Y |
| JAZ-016 |  | 503 |  | 30,265 |  | N |  |  | Y |

* One data point without a measured absorbance

4.0	Quality Assurance and Quality Control

This section presents the QA/QC procedures that were followed in order to ensure the highest possible data quality.

4.1	Time Synchronization

The internal time for each instrument was set by hand on the day of each trial.  For Trial 1, this calibration was +/- 60 seconds using an external clock (mobile phone) not confirmed to be synchronized with the official DPG (JR2 command center) timekeeping system.  After conferring with program staff, instruments were calibrated to +/- 1 second for Trials 2-5 using the official DPG time.  Calibration within a one second interval was achieved utilizing a spotter calling the time to  the person setting up the instrument. Calibration to <1 second intervals is not possible with the current JazTM UV-Vis DPU.

4.2	Instrument Calibration

SigSci staff performed pre- and post-trial QC checks on each JazTM UV-Vis instrument for each trial. These calibration checks were performed using a ppm chlorine standard in a Tedlar® bag.  Readings for each properly functioning instrument showed high concordance between pre- and post-trial readings (see figures 2 and 4-7).

SigSci staff also performed calibration and evaluation checks on each JazTM UV-Vis instrument prior to delivery to DPG. This testing utilized chlorine standards of 0, 100, 1,000, 10,000 and 100,000 ppm. R2 values from the resulting linear calibration curves fell between 0.9977 and 1 for all instruments (Appendix A).

4.3	Sampling Rate

The JazTM UV-Vis instruments were set to record absorbance readings every two seconds. However, the exact frequency of these measurements varied due to the write speed of the individual instrument’s DPU to the SD card. Overall, the majority of instruments sampled, on average, within one hundredth of a second of this interval (see table 7). However, three instruments (specifically 002, 005, and 015) consistently sampled more slowly than expected. Extended data recording times were likely due to slower write speeds on these instruments. While the temporal resolution provided by these three instruments is slightly lower as compared to the others, a significant amount of data was still collected over the period of each release.

Table 7: Sampling Rate Data per Instrument for All Trials

| Instrument | Trial 1 | Trial 2 | Trial 3 | Trial 4 | Trial 5 | Average |
| --- | --- | --- | --- | --- | --- | --- |
| JAZ-001 | 2.008 | DNC | 2.007 | 2.007 | 2.015 | 2.009 |
| JAZ-002 | 2.339 | 2.345 | 2.303 | 2.395 | 2.471 | 2.371 |
| JAZ-003 | 2.008 | 2.012 | 2.000 | 2.000 | 2.000 | 2.004 |
| JAZ-004 | 2.004 | 2.004 | 2.004 | 2.000 | 2.004 | 2.003 |
| JAZ-005 | 2.257 | 2.237 | 2.109 | 2.105 | 2.081 | 2.158 |
| JAZ-006 | 2.008 | DNC* | 2.004 | 2.000 | DNC | 2.004 |
| JAZ-007 | 2.016 | DNC | 2.004 | 2.000 | 2.004 | 2.006 |
| JAZ-008 | 2.012 | 2.008 | 2.000 | 2.007 | 2.019 | 2.009 |
| JAZ-009 | 2.008 | 2.008 | 2.007 | 2.000 | 2.007 | 2.006 |
| JAZ-010 | DNC | DNC | DNC | 2.011 | 2.019 | 2.015 |
| JAZ-011 | 2.012 | 2.008 | 2.011 | 2.011 | 2.004 | 2.009 |
| JAZ-012 | DNC | DNC | 2.011 | 2.011 | 2.007 | 2.010 |
| JAZ-013 | 2.016 | 2.012 | 2.007 | 2.004 | 2.007 | 2.009 |
| JAZ-014 | 2.008 | 2.008 | 2.004 | 2.007 | 2.011 | 2.008 |
| JAZ-015 | DNC | DNC | 2.438 | 2.308 | 2.343 | 2.363 |
| JAZ-016 | 2.004 | 2.016 | 2.007 | 2.007 | 2.004 | 2.008 |

*DNC: did not collect

4.4	Environmental Data

Environmental conditions (temperature, humidity, barometric pressure) were not recorded or considered as part of this report. Meteorological data for each trial was collected separately by DPG to enable any necessary corrections for environmental conditions.

4.5	Field Test Data

Data from each trial was reviewed immediately following each trial to survey for consistency and anomalies. These preliminary data summaries were delivered to Dr. Shannon Fox and trial participants who have government email addresses within 48 hour after a trial via the AMRDEC website.  The summaries consisted of a word document and a spreadsheet containing data for each instrument.  As mentioned in section 2, the text files with the instrument raw data files with instrument locations for each trial were delivered to Mr. Kendal Ferguson at DPG by AMRDEC.  When possible, instruments which failed to record data or provided anomalous readings were repaired prior to subsequent trials.

Background levels for each instrument were tracked before and after each trial. Baseline values after each trial consistently returned to within one standard deviation of the mean baseline recorded in the minute preceding the chlorine release.

4.6	Data Storage and Retrieval

Data was processed electronically by the on-board JazTM UV-Vis software prior to storage on an SD card for later retrieval. Following each field trial, data was retrieved from the SD card onto the hard drive of a laptop computer and then written to thumb drive as a backup. Only after storing the data on two different media were the raw data files erased from the SD card.

5.0	Recommendations

The JazTM UV-Vis instruments performed well throughout Jack Rabbit II testing and present perhaps the most reliable method to measure chlorine concentrations up to and exceeding 100,000 ppm. However, several shortcomings have been identified as part of both Jack Rabbit I & II testing which, if addressed, would improve data quality, data interpretation, and ease-of-use.

5.1	Standard Curve

It was previously understood that chlorine concentrations significantly greater than 100,000 ppm (corresponding to approximate absorbance values greater than 2.5) were considered saturating for the JazTM UV-Vis instruments.  However, after reviewing the data surrounding these trials, only one instrument encountered a definitive saturation event wherein the absorbance during the highest concentration portion of the release could not be recorded. As the highest standard used to generate the standard curve for these trials was 100,000 ppm, the accuracy surrounding the measurement of higher values cannot be confirmed. SigSci recommends that higher concentration standards (200,000-500,000 ppm) are used in the future to establish the upper limit of detection for these instruments. These concentrations will generate additional health and safety concerns, however, and will need to be confined to an appropriate controlled laboratory.

In addition, adding a 50,000 ppm standard to the existing curve will provide better fidelity to the existing curve.  A 50,000 ppm standard can be used in the current laboratory set up and will not require additional safety considerations beyond those already in place for the 100,000 ppm standard.

5.2	Time Synchronization

Time synchronization remains a hurdle for the JazTM UV-Vis instruments. Despite a processor upgrade prior to Jack Rabbit II, some instruments experienced variable write-times to the SD card which decreased the temporal resolution of the instruments (Table 7). In the future, fitting each instrument with a hard drive or memory device with a much faster write speed will ensure that consistent data is captured across the trial. Time synchronization and data retrieval are also complicated by the inability of the JazTM UV-Vis instruments to be wirelessly networked. Time must be manually entered as part of the setup process for each instrument. While the instrument measures time to within one one-thousandth of a second, manual input is generally only accurate to within +/- one second. Future upgrades to allow the instruments to synchronize their internal clocks and transmit data to a central computer will improve data quality and decrease the amount of labor associated with field technician-dependent instrument maintenance and setup.

Appendix A: Calibration Plots

Appendix B: Individual Instrument Plots

Appendix C: Specific Instrument Locations per Trial

Table 8: Specific Instrument Locations per Trial

| Trial 1 |  |  |  |
| --- | --- | --- | --- |
| JAZ Unit | CONEX | Location Description | Survey Number |
| JAZ-001 | 6.2 | Downwind of CONEX | NA |
| JAZ-002 | 7.3 | Downwind of CONEX | NA |
| JAZ-003 | 7.1 | Downwind of CONEX | NA |
| JAZ-004 | 7.5 | Downwind of CONEX | NA |
| JAZ-005 | 8.5 | Downwind of CONEX | NA |
| JAZ-006 | 2.1 | Downwind, left corner | NA |
| JAZ-007 | 6.8 | Downwind of CONEX | NA |
| JAZ-008 | 11.4 | Inlet to CONEX | NA |
| JAZ-009 | 1.4 | Downwind of CONEX | NA |
| JAZ-010 | 6.5 | Downwind of CONEX | NA |
| JAZ-011 | 1.2 | Downwind of CONEX | NA |
| JAZ-012 | 1.3 | Downwind of CONEX | NA |
| JAZ-013 | 9.4 | Downwind of CONEX | NA |
| JAZ-014 | 11.5 | Inlet to CONEX | NA |
| JAZ-015 | 11.3 | Inlet to CONEX | NA |
| JAZ-016 | 11.2 | Inlet to CONEX | NA |
| Trial 2 |  |  |  |
| JAZ Unit | CONEX | Location Description | Survey Number |
| JAZ-001 | 9.1 | Downwind of CONEX | J91 |
| JAZ-002 | 11.4 | Downwind of CONEX | NA |
| JAZ-003 | 7.1 | Downwind of CONEX | J71 |
| JAZ-004 | 7.5 | Downwind of CONEX | J75 |
| JAZ-005 | 8.5 | Downwind of CONEX | J85 |
| JAZ-006 | NA | Upwind, NW corner of apron | P10 |
| JAZ-007 | 9.7 | Downwind of CONEX | J97 |
| JAZ-008 | 11.5 | Inlet to CONEX, right side | J115 |
| JAZ-009 | 1.5 | Downwind of CONEX | J15 |
| JAZ-010 | NA | Out of service | P4 |
| JAZ-011 | 1.1 | Downwind of CONEX | J11 |
| JAZ-012 | NA | Upwind NE corner of apron | P13 |
| JAZ-013 | 9.4 | Downwind of CONEX | J94 |
| JAZ-014 | 11.5 | Inlet to CONEX, left side | J115 |
| JAZ-015 | 11.3 | Inlet to CONEX, left side | J113 |
| JAZ-016 | 11.3 | Inlet to CONEX, right side | J113 |
| Trial 3 |  |  |  |
| JAZ Unit | CONEX | Location Description | Survey Number |
| JAZ-001 | 9.1 | Downwind of CONEX | J91 |
| JAZ-002 | 11.4 | Downwind of CONEX | NA |
| JAZ-003 | 7.1 | Downwind of CONEX | J71 |
| JAZ-004 | 7.5 | Downwind of CONEX | J75 |
| JAZ-005 | 8.5 | Downwind of CONEX | J85 |
| JAZ-006 | NA | Upwind NW corner of apron | P10 |
| JAZ-007 | 9.7 | Downwind of CONEX | J97 |
| JAZ-008 | 11.5 | Inlet to CONEX, right side | J115 |
| JAZ-009 | 1.5 | Downwind of CONEX | J15 |
| JAZ-010 | NA | Downwind off apron | P4 |
| JAZ-011 | 1.1 | Downwind of CONEX | J11 |
| JAZ-012 | NA | Upwind NE corner of apron | P13 |
| JAZ-013 | 9.4 | Downwind of CONEX | J94 |
| JAZ-014 | 11.5 | Inlet to CONEX, left side | J115 |
| JAZ-015 | 11.3 | Inlet to CONEX, left side | J113 |
| JAZ-016 | 11.3 | Inlet to CONEX, right side | J113 |
| Trial 4 |  |  |  |
| JAZ Unit | CONEX | Location Description | Survey Number |
| JAZ-001 | 9.1 | Downwind of CONEX | J91 |
| JAZ-002 | 11.4 | Downwind of CONEX | NA |
| JAZ-003 | 7.1 | Downwind of CONEX | J71 |
| JAZ-004 | 7.5 | Downwind of CONEX | J75 |
| JAZ-005 | 8.5 | Downwind of CONEX | J85 |
| JAZ-006 | NA | Upwind, NW corner of apron | P10 |
| JAZ-007 | 9.7 | Downwind of CONEX | J97 |
| JAZ-008 | 11.5 | Inlet to CONEX, right side | J115 |
| JAZ-009 | 1.5 | Downwind of CONEX | J15 |
| JAZ-010 | NA | Downwind off Apron | P4 |
| JAZ-011 | 1.1 | Downwind of CONEX | J11 |
| JAZ-012 | NA | Up wind, NE corner of apron next to scrubber | P13 |
| JAZ-013 | 9.4 | Downwind of CONEX | J94 |
| JAZ-014 | 11.5 | Inlet to CONEX, left side | J115 |
| JAZ-015 | 11.3 | Inlet to CONEX, left side | J113 |
| JAZ-016 | 11.3 | Inlet to CONEX, right side | J113 |
| Trial 5 |  |  |  |
| JAZ Unit | CONEX | Location Description | Survey Number |
| JAZ-001 | 9.1 | Downwind of CONEX | J91 |
| JAZ-002 | 11.4 | Downwind of and against CONEX, at junction between the bottom two of the stack | NA |
| JAZ-003 | 7.1 | Downwind of CONEX | J71 |
| JAZ-004 | 7.5 | Downwind of CONEX | J75 |
| JAZ-005 | 8.5 | Downwind of CONEX | J85 |
| JAZ-006 | NA | Upwind, NW corner of apron | P10 |
| JAZ-007 | 9.7 | Downwind of CONEX | J97 |
| JAZ-008 | 11.5 | Inlet to CONEX, right side | J115 |
| JAZ-009 | 1.5 | Downwind of CONEX | J15 |
| JAZ-010 | NA | Downwind off Apron | P4 |
| JAZ-011 | 1.1 | Downwind of CONEX | J11 |
| JAZ-012 | NA | Up wind, NE corner of apron next to scrubber | P13 |
| JAZ-013 | 9.4 | Downwind of CONEX | J94 |
| JAZ-014 | 11.5 | Inlet to CONEX, left side | J115 |
| JAZ-015 | 11.3 | Inlet to CONEX, left side | J113 |
| JAZ-016 | 11.3 | Inlet to CONEX, right side | J113 |

Appendix D: List of Acronyms

| Acronym | Full Name |
| --- | --- |
| AMRDEC | Army Missile Research Development and Engineering Center |
| CONEX | Shipping Container |
| CSAC | Chemical Security Analysis Center |
| DHS | Department of Homeland Security |
| DNC | Did Not Collect |
| DPG | Dugway Proving Grounds |
| DPU | Data Processing Unit |
| JR 2 | Jack Rabbit 2 |
| LBL | Lawrence Berkley National Laboratory |
| QA | Quality Assurance |
| QC | Quality Control |
| PPM | Part Per Million |
| SD | Secure Data |
| SigSci | Signature Science, LLC |
| TIC | Toxic Industrial Chemical |
| TIH | Toxic Inhalation Hazard |